## Slide 1: Computed Radiographic Imaging of Ceramic Body Armour Personal Protective Equipment
Muzibur Khan, Michael Brothers, Trent Gillis
Non-Destructive Evaluation, Aerospace Research Centre
Notes: Good Afternoon. Welcome to my presentation on “Computed Radiographic Imaging of Ceramic Body Armour Personal Protective Equipment”.
My name is Muzibur Khan. I am working as Research Office in Non-Destructive Evaluation Group, Aeropsace Research Centre of National Research Council of Canada
Lets introduce my colleagues who also co-authors of this presentation, paper and project team member
Michael Brothers, Trent Gillis
My presentation primarily will focus on the expereince and lesson learned from the application of CR for detecting damages in significant numbe rof PPE plates made from composite vceramic materials
It presents the demonstration and validation of CR based PPE inspection processes, along with CR process control mechanism to ensure optimum image quality, high repeatability, while maintain a low false positive call rate.
You can find additional information about this presentation in the paper submitted for publication in the CINDE proceeding

## Slide 2: Background
Ballistic protection, integrity and extended service life are key requirements of PPE Plates
Ballistic performance of PPE are degraded due to presence of damage (in-service manufacturing)
Presence of damage compromises the required protection
NDE required for rapid determination of presence of damage, suitability of PPE plates (protective) for combat deployment.
Radiographic NDE capable to visualize/quantify the presence of damage, verifying the integrity of PPE plates (high throughput, insensitive to variations in the PPE)
Notes: Here little background of this project
Body armor PPE has been in use for a very long time and is generally related to bullet proof vest worn for ballistic protection to important body parts.
Ballistic performance of ceramic body armour plates (PPE) are degraded due to presence of damage (in-service manufacturing)
Currently there is no well understood mechanism or accepted practice for determination life expectancy of body armour and their life span varies significantly.
it is important to ensure that the structural integrity of the armor is maintained throughout its life cycle, preserved through periodic inspection.
High cost (manufacturing and raw materials) of ceramic armours has created the expectation of extended service life while maintaining the integrity of personal protective equipment PPE plates and performance during the period of service.
Structural complexity and signal dispersion in armour composite materials with many interfaces, including composite materials, can make ultrasonic inspection very difficult
Radiographic Non-destructive evaluation (NDE) techniques are required to rapidly and non-destructively determine the suitability of ceramic protective plates for combat deployment.
CR allows for fast, reliable and accountable inspections. Film radiography also lack digital advantages such as archiving, enhancement, filter, etc.

## Slide 3: Computed Radiography (CR)
CR uses a reusable phosphor imaging plate and allows faster image acquisition digitally, improved capabilities for analysis and digital storage
Notes: This pictorial demonstrating how the CR system works. Phosphorous plate is loaded into a cassette and is exposed to an X-ray or gamma ray radiation source. During the radiation exposure, the absorbed radiation energy is stored in the phosphorous plate and forms a latent image of the object. The exposed imaging plate is run through a laser reading device that scans the plate. As the laser beam scans the surface of the phosphorous imaging plate in a raster pattern, light is emitted where radiation energy was previously stored. The intensity of light emitted from the IP is proportional to the amount of radiation absorbed by the storage phosphor.

## Slide 4: Computed Radiography at NRC
Computed Radiographic System (SR 70 µm, HR 35 µm)
  16 Bit, Scan width 14”, Throughput 90 plates/hr (SR), 28 plates/hr (HR)
CR Image Acquisition/Analysis Software
HR & medium speed (IPS) and extreme HR and slow speed imaging plates (IPU)
BAM Certified highest system Class IP 1 / 80 according to ISO 16371-1 and ASTM E2446, ASTM DICONDE Compliant, GE CR Phantom for System/Image QC
Notes: High performance table-top flatbed scanner with contactless plate transport, inline erasure
IPS High Resolution & Medium Speed (for inspection of welds), IPU (Extreme High Resolution & Slow Speed - for extremely high resolution applications when very low micron range sensitivity is required)
Standard resolution of 70 µm, and high resolution of 35 µm

## Slide 5: Primary Metrics for Image Quality in CR
Primary metrics for establishing the performance are
  Signal-to-Noise Ratio (SNR)
  Contrast-to-Noise Ratio (CNR)
  Basic Spatial Resolution (SRb)
  Higher the SNR, CNR or smaller the SRb, better the quality
  In film radiography (RT) these are evaluated indirectly by image quality indicators (IQI)
  No direct relationship between radiographic sensitivity and flaw size
    SNR 226.273
    (CNR) < 2
    (SRb)
Notes: Image quality and degree of details that a CR image can reveal are the key performance indicators considered during characteristics evaluation.
For computed radiography (CR), the primary metrics for establishing the performance are signal-to-noise ratio (SNR), contrast-to-noise ratio (CNR), and basic spatial resolution (SRb)
Signal term is used for the mean pixel/gray values of a digital image.
Contrast term is used to quantify the difference of the radiation signal intensity (measured as pixel value) between the base material and the discontinuity (replicated using an IQI)
Basic Spatial Resolution or sharpness of image (SRb) is the ability to distinguish between small objects that are close together in digital image or ability to clearly see abrupt changes in or ability to clearly see abrupt changes in an object
SNR is the ratio of the mean pixel values (signal) to the standard deviation of the pixel values (noise).
CNR is defined as the ratio of the difference of signal intensities of two regions of interest to the background noise (standard deviation of the pixel values)
In film radiography applications, these metrics are typically evaluated indirectly using an image quality indicator in a specified film density range.

## Slide 6: Contrast in CR
In general CR has better “latitude” than film
CR can cover a wide range of thicknesses/densities in a single exposure
CR has better contrast sensitivity (CS)
  i.e. Capability to “window/level” to visualize a contrast step
CR
film
1%
2%
3%
4%
Section 1 – The Image Plate
Film Latitude
Film
Digital Latitude
CR
Exposure
Pixel value
or
film density
Ref: Digital and Film Radiography Comparison - Ken LaCivita
Notes: In general CR has better latitude than film.

## Slide 7: Inspection Specimen (PPE Body Armour)
10 inches by 12 inches in size (Approx.)
Hard front face of ceramic dimple (front face)
High performance polymer composite (back face)
Sensitive material in terms of their fabrication and intellectual property, preventing to provides additional details info
Damage may formed during manufacturing and in-service operational condition
Notes: Here is the pictorial of a typical PPE plates used in the inspection. Its about 10 inches by 12 inches.
Hard front face of ceramic dimple (front face)
High performance polymer composite (back face)
PPE materials are sensitive material in terms of their fabrication and intellectual property, preventing to provides additional details info
Different kind of damages may formed during manufacturing and in-service operational condition
Digital computed radiographic techniques was deployed for several case studies to detect damages in PPEs

## Slide 8: Specimen Setup
Notes: Source to detection (IP) distance were maintained at 60 inches
The exposure (X-ray Voltage, current and exposure time) also adjusted to achieve target grey scale value range on the center dimples of the plates, grey-scale values on the edge portion of the plates and to avoid over exposure which may lead to false rejection
In this study the CR digital images were acquired in universally compatible Digital Imaging and Communication in Nondestructive Evaluation (DICONDE) format and the detected damages/anomalies and their characteristics were tagged in the digital images

## Slide 9: Image Quality Performance Indicator
Al-2024 (Aluminum alloy 2024) step-wedge block
To ensure optimum image quality
Repeatability of results among various CR systems in accordance with the dynamic range
exposure setting aimed to yield dynamic range (pixel or grey scale) distributed from 10% to 90% of the maximum pixel value.
Two ASTM shim blocks to verify image quality meet 2% contrast sensitivity
Notes: Al-2024 (Aluminum alloy 2024) step-wedge block were placed in the side of specimen and imaging plate.
To ensure optimum image quality and repeatability of results among various CR systems in accordance with the dynamic range of the individual CR system
The exposure setting aimed to yield dynamic range (pixel or grey scale) distributed from 10% to 90% of the maximum pixel value.
Two ASTM image quality indicators (IQI) shim blocks (0.087 inches and 0.062 inches) were placed on top of two steps (0.620 inches, 0.087 inches) to verify that the achieved image quality meet 2% contrast sensitivity, which commonly required contrast in most engineering codes in CR process control.
For correct exposure setting, all steps on the Al-2024 step-wedge, with the exception of the two thinnest steps, are required to be visible and equally discernable on the CR image.

## Slide 10: Sample Found Defects (Cracks)
Cracks of different size (Few millimetres to the through length or width dimension of ceramic plates)
probable causes of cracks: pressure of packaging during pressure-cast forming manufacturing stage
Relieve local stress from external forces (dropped or knocked)
Notes: Randomness of crack sizes ranged from few millimetres to the through length or width dimension of the ceramic plates

## Slide 11: Sample Found Defects (Void)
Probable cause of edge void flaws is chipped edge of the ceramic due to handling (service or manufacturing)
Voids at the armour centre dimple area likely caused due to chipped dimple on the surface of armour plate
Notes: Void flaws which is another cause for rejection also found in the plates.

## Slide 12: Sample Found Defects (Miro-tears /Micro crack)
Notes: micro-cracks and/or micro-tears were also noticed in the armor plates.

## Slide 13: Sample Found Defects (Porosity)
Notes: During the inspection inherent manufacturing process induced anomaly or discontinuity indications such as cluster of porosity,

## Slide 14: Found Defects (Missing dimple at edge)
Notes: The plates were accepted or rejected based on pass/fail criteria which was established based on prior ballistic performance test. The underlying reason behind such rejection is that plate’s total integrity or ballistic performance is compromised because of that kind of defect.
However, due to sensitivity of the data and performance test results, the criteria is restricted to publish.

## Slide 15: Results & Conclusions
Significant number of PPE plates were inspected, damage types were characterized and evaluated
Cracks of varying size and void-like discontinuities in accordance with applicable inspection criteria were the primary causes for rejection of plates.
Standard resolution imaging successful were successful in revealing manufacturing inherent and service induced non-critical flaws
Computed radiography demonstrated its capability to provide high throughput in inspection, unique relative insensitivity to the considerable variations
To further ease the inspection process, we plan to apply machine learning/artificial intelligence approach to aid in automatic or semi-automatic damage detection/classification
Notes: Service induced cracks and voids were primary sources of defects for plate rejection. The cracks profile showed the randomness of crack sizes ranged from few millimetres to the through length or width dimension of the ceramic plates.
In addition to cracks and voids CR imaging using standard resolution imaging successful were successful in revealing manufacturing inherent and service induced non-critical flaws include porosity, micro-tears or micro-cracks, missing dimples at the edge, trimmed dimple at the edge, small voids at edge (<1 cm-diameter), ceramic particle through-out, gross edge thinning voids etc.

## Slide 16: THANK YOU
Muzibur Khan • Research Officer • muzibur.khan@nrc-cnrc.gc.ca